## Slide 1: Optics of Ophthalmic Instruments
Haleh Ebrahimi, OD
Clinical Assistant Professor
University of Pittsburgh Medical Center

## Slide 2: Most used instruments
Binocular Indirect Ophthalmoscope
Retinoscope
Lensometer
Keratometer

## Slide 3: Binocular Indirect Opthalmoscope

## Slide 4: Principles of a Slide Projector
  An intense source of light illuminates a transparent slide from behind and makes that slide a Luminous object on its own
  That luminous object emits light which is picked up by the projector lens

## Slide 5
Principles of a Slide Projector
  The lens is situated so that the object is just beyond its primary focal plane
  The final image can then be projected onto a screen

## Slide 6: Binocular Indirect Ophthalmoscope
With the indirect, you make the fundus a luminous object (the slide of the last example)
The eye acts like a miniature opaque projector and projects that object out into space
This is done by shining an intense source of light through the pupil – the fundus image is then projected out

## Slide 7
An emmetropic eye would project that image to infinity
We use a 20D hand held condensing lens to create the image in its secondary focal plane
This image is called an aerial image and is inverted when compared to the orientation of the fundus itself

## Slide 8: Optics of BIO
That aerial image has thickness representing any thickness present in the fundus itself
Any excavation of the nerve head or elevated of a choroidal lesion will be manifested in the “depth” of the image
The longer the focal length of the condensing lens, the farther from the lens the image will be  increased lateral magnification

## Slide 9: Question
Which of the following lenses provided the largest field of view?
  15 D lens
  20 D lens
  30 D lens
  Field of view is the same in all the lenses

## Slide 10: Question
Which of the following lenses provided the largest field of view?
  15 D lens
  20 D lens
  30 D lens
  Field of view is the same in all the lenses
  Field of view = dioptric power x 2

## Slide 11: Binocular Indirect Ophthalmoscope
Axial images dimensions are especially THICKENED by the LOWER powered (longer focal length) condensing lens

## Slide 12: Binocular Indirect Ophthalmoscope
30 Diopter Lens (highest magnification)
  Yields LEAST magnification of retina 60/30=2x
  Stereopsis is half that of the normal 2/4=1/2
  Largest field of view 30x2= 60 degrees
15 Diopter Lens
  Magnification of retina 60/15= 4x
  Stereopsis 4/4=1  FULL STEREOPSIS = thickened image
  Field of view 15x2= 30 degrees
20 D lens is most used since it provides adequate field of view, stereopsis and magnification

## Slide 13: Advantages of Binocular Indirect Ophthalmoscopy
Stereoscopic visualization
Increased field of view vs. direct ophthalmoscopy (25° vs 8°)

## Slide 14: Light Reflexes
3 bright reflexes (from the cornea, pupil and crystalline lens) when trying to visualize the fundus
The condenser is a biconvex lens, the anterior surface forms a virtual reflected image of the illuminating headlamp, and the posterior surface, being concave to the illuminating source, forms a real image
These images move in opposite directions if the lens is tilted
The higher the D.P. Of the condensing lens, the more will be tilting necessary to move reflexes away from the center of view

## Slide 15: Board Question

## Slide 17: Retinoscope
To determine the refractive error of an eye objectively
It is a small projector which emits a spot or streak like image of the lamp filament itself
This projector is aimed at the patient’s eye

## Slide 18: Retinoscope
The retinoscopist can project that image onto a wall any distance away and focus it sharply, from infinity to any place behind or in front of the patient’s eye
2 principals:
  Plano-mirror Effect
  Concave-mirror Effect

## Slide 19: Plano-Mirror Effect
Effective light source lies behind the plane of the mirror (most common used)
The rays of light from the source goes parallel or slightly diverging
Does not cross between the source and the patient’s eye
  With movement – hyperopia
  Against movement – myopia
Direction of motion of the light source on the retina will always be in the same direction as the motion of the retinoscope light as it crosses along the patient’s face

## Slide 21: Concave-Mirror Effect
Keep the effective source in front of the plane of the mirror
rays emitted from source are more converging
cross at a certain distance between patient and the source
    With movement – Myopia
    Against Movement – Hyperopia
Not used as often
Direction of motion of the light source on the retina will always be in the opposite direction as the motion of the retinoscope streak

## Slide 23: Retinoscopy
The closer the observer’s eye the patient’s luminous retinal image is located, the larger and brighter will be the illuminated shadow and the quicker it will move
Plus lens used to counter any observed “with” motion
Minus lens for “against” motion
Goal is to “Neutralize” the observed motion with the lens

## Slide 24: Working Distance
To find the “corrective” lens to help the patient see optical infinity, we must image the retina not at the retinoscope but further back behind the observer, at infinity
The lens has to be a power of 1/d
  Ex. 50 cm working distance
    1/0.5=2
    Means subtract 2D from the lens you found to neutralize the reflex
For an astigmatic eye, each meridian must be neutralized separately

## Slide 25: Board Question
A. +1.00 + 2.75 x 135
B. +1.00 + 2.75 x 045
C. +3.00 + 2.75 x 135
D. +3.00 +2.75 x 045

## Slide 27: The Lensometer

## Slide 28: Lensometer
To quickly determine the power, cylinder and axis of an “unknown” lens
Can measure the amount and direction of any prism that may be incorporated or induced by the lens
Most clinically important information about any corrective lens is the distance between its back surface and its secondary focal point = back focal length

## Slide 29: A Simple Lensometer
The vertex power (effective power) = the reciprocal of the back focal length

## Slide 30: A Simple Lensometer
A grid like, luminous target is moved back and forth until light rays leaving the lens are parallel
  (as detected by a small eyepiece telescope which is focused for infinity so that only the parallel light bundles will appear to be clear)
The target must then be located at F’ of the lens
Its distance u is simply measured and converted to diopters
This method is impractical because the instrument would have to be too long

## Slide 31: A Lensometer
Uses Badal’s Principle to shorten the length needed to measure the power of an unknown lens
The fixed field lens is placed so that its focal point exactly coincides with the posterior vertex of the unknown lens
The observer moves the illuminated target back and forth in relation to the fixed field lens until the reticle target appears sharply in focus through the eyepiece telescope

## Slide 33: The Lensometer
With the focal point of the field lens situated exactly at the posterior vertex of the “unknown” lens, the proper optical condition is set up so…
  The distance u is DIRECTLY proportional to the power of the unknown lens
The scale will not shrink as higher power lenses are tested

## Slide 34: The Badal Principle
Knapp’s Law applied to Lensometers

## Slide 35: The Badal Principal
Place any plus lens in front of the eye so that the lens’ F’ coincides with the eye’s nodal point
An object is put in front of that lens, but closer than its primary focal point F

## Slide 36: The Badal Principal
The object tip will emit one ray which is parallel to the axis
After refraction, that ray will pass through F’
Since F’ was made to coincide with the nodal point of the eye, the ray will continue undeviated

## Slide 37: Determining an Add with a Lensometer
To obtain the distance power of any lens, you must use the Lensometer to measure the back focal length
The front focal length is a better indicator of the ability of a “thick lens” to deal with near objects
If the add on the aphakic lens is read with the lens in the ‘normal’ instead of ‘reversed’ position, the indicated power will be too high in plus compared to its actual effectivity

## Slide 38: Board Question
A light source falls on an unknown lens. A line is formed at 45° at 20 cm. When the screen is moved to 33cm, the line is at 135°. At what distance from the lens is a circle seen on the screen?
  A. 20cm
  B. 22.5 cm
  C. 25 cm
  D. 26.5 cm

## Slide 40: Board Question
A light source falls on an unknown lens. . A line is formed at 45° at 20 cm. When the screen is moved to 33cm, the line is at 135°. What is the prescription of the lens?
A. +3.00 +2.00 x 045
B. +4.00 Sphere
C. +3.00 + 5.00 x 045
D. +5.00 + 3.00 x 135

## Slide 42: Keratometer

## Slide 43: The Keratometer
To measure the corneal curvature
Measures the anterior 3mm surface of the cornea
Used mainly for contact lens fitting
Keratometer measures the size of the reflected image to give the radius of curvature of that anterior corneal surface

## Slide 44: Information provided by a Keratometer
The radii of curvature of the cornea
The directions of the principal meridians of the eye, determining whether the astigmatism is with-the-rule or against-the-rule
  a key factor in contact lens fitting
The degree of corneal astigmatism
The presence of any corneal distortion (keratoconus, pterygium, corneal scarring)

## Slide 45: Optical Principles
Cornea is a convex refracting surface
To find the refracting power of the cornea, we need to reflect an object of a known size at a known distance off the corneal surface.
Then determine the size of the reflecting image with measuring telescope and calculate the refractive power of the cornea based on the refractive index of n=1.3375

## Slide 48: Keratometer
The radius of curvature of of a mirror will determine the size of the image it produces from a given sized object
If we illuminate an object of known size, place it a known distance from the cornea, are are able to measure the size of the reflected image, we can deduce the radius of curvature
The cornea is considered a high powered mirror (-250 D), an object does not have to be too far away to be effectively at the cornea’s optical infinity

## Slide 51: Surface Power Formula
The keratometer usually shows the curvature in diopters of power or in millimeters and diopters
If a millimeter scale is not given, it can be found from tables or by using the surface power formula using the refractive index 1.3375
Surface Power Formula:
  D = the dioptric power of the Cornea
  n = the refractive index of the keratometer used (1.3375)
  R = the radius of curvature of the cornea in meters
Notes: https://www.slideshare.net/hmirzaeee/keratometry-31466440

## Slide 53
Notes: https://www.slideshare.net/KaranBhatia25/keratometry-55540991

## Slide 55: Question
A standard manual keratometer measures central corneal power at what approximate diameter?
1.65
2.5
3.2
5.0